# Applicability determination — SAMPLE

Illustrative output for a sample product. Not advice for any real product.

| Item | Detail |
| --- | --- |
| Product | Sample: automated upper-arm blood-pressure monitor |
| Description | Reusable, non-invasive, battery-powered oscillometric monitor with an inflatable arm cuff and a companion app over Bluetooth. Home and clinical use, adults. Sold in the US and EU. |
| EU classification | Class IIa (Annex VIII, Rule 10 — active device for diagnosis/monitoring) — Regulation (EU) 2017/745 (MDR) |
| US classification | Class II, 510(k) — 21 CFR 870.1130 — non-invasive blood pressure measurement system |

## Standards

| Standard | Title | Why it applies |
| --- | --- | --- |
| IEC 60601-1 | Medical electrical equipment — basic safety and essential performance | Electrically powered device |
| IEC 60601-1-2 | Electromagnetic disturbances (EMC) | Electronics + Bluetooth radio |
| IEC 60601-1-11 | Home healthcare environment | Intended for home use |
| IEC 80601-2-30 | Automated non-invasive sphygmomanometers | Particular standard for this device type |
| ISO 81060-2 | Clinical investigation of intermittent automated measurement type | Clinical validation of accuracy |
| ISO 14971 | Application of risk management to medical devices | All medical devices |
| IEC 62304 | Medical device software — life cycle processes | Embedded firmware + app |
| IEC 62366-1 | Usability engineering | Lay users at home |
| ISO 10993-1 | Biological evaluation — evaluation and testing | Cuff in contact with intact skin |
| ISO 15223-1 | Symbols for labels and information | Labeling |

Sample generated by the public demo. In a workspace, every item carries confidence, source and a named reviewer's decision.